## Slide 1: Microfluidic Background Information

## Slide 2: What is Microfluidics?
The application of science & technology to manipulate tiny fluid volume (nanolitres-microlitres) in microscale channels
Complex lab functions can be scaled down and simplified
  Known as “lab-on-a-chip”
Characteristic Physics:
  Predominantly laminar (non-turbulent) flow)
  Precise temporal and spatial control of fluid movement
Used in physics, chemistry, biology and engineering applications.

## Slide 3: Benefits
Reduces sample and reagent consumption (cost)
Point-of-use application
Integrates multiple laboratory steps on a chip
Rapid development of process line and fast analysis

## Slide 4: Main Application Fields I
Biomedical and Life Sciences
  Point-of-care (at the time and place of a patient care/in the moment) diagnostics
    E.g. for blood or infectious diseases
  DNA/RNA, protein analysis and PCR tests
  Single cell studies and cell manipulation
  Organ-on-a-chip models
Pharmaceutical Sciences
  Drug screening and formulation
  Synthesis of nanoparticles and lipid carriers

## Slide 5: Main Application II
Environmental and Food Monitoring
  Detection of pollutants, contaminants
  Agricultural biosensing
Chemical Analysis and Synthesis
  Miniaturised reactors and rapid chemical synthesis
  Reaction optimisation
Industrial/Basic Research
  Synthesis of micro/nanoparticles
  Studies of fluid physics at the microscale

## Slide 6: Presentation Contents
Microfluidic Focus:
  Point-of-care (POC) devices
  In-vitro diagnostics (IVD)
  Lipid nanoparticles (LNP) mixing applications
content:
  Chip design
  Materials
  Manufacturing/Fabrication
  Roll-to-Roll vs. Injection Moulding
  Recommendations

## Slide 7: Application fields of Microfluidic Chips
POC Diagnostics
  Rapid, on-site disease testing (e.g. HIV, COVID-19) compared to lab equipment
  Small sample volumes (20-30 µL), fast results (10-30 min)
  Portable, often single use
  Detection Methods: Optical/Molecular/Smartphone
  Multipurpose PoC can identify multiple diseases using electronic devices
Notes: POC
https://pmc.ncbi.nlm.nih.gov/articles/PMC8875995/
https://www.ufluidix.com/microfluidics-applications/microfluidic-devices-for-point-of-care/
https://www.elveflow.com/blog/transforming-healthcare-microfluidic-chips-in-point-of-care-diagnostics
IVD

## Slide 8: Application fields of Microfluidic Chips
IVD (In Vitro Diagnostics)
  Automated, compact molecular, protein and genetic analysis
    “Miniturising complex assays on low cost platforms”
  Multiplexed, high-throughput for microfluidics
  Lab-quality results in closed, safe system … difficult to scale up
Notes: IVD
https://en.seamaty.com/index.php?s=/sys/483.html
https://www.creative-biolabs.com/drug-discovery/diagnostics/microfluidics.htm
https://pmc.ncbi.nlm.nih.gov/articles/PMC9469643/

## Slide 9: Application fields of Microfluidic Chips
LNP Mixing (Pharma/Biotech)
  LNP’s are encapsulated lipids (e.g. mRNA), used as nano-delivery vehicles for cytotoxic chemotherapy/antibiotics and nucleid acid therapeutics
    = Drug delivery
  Controlled mixing via (chaotic) microstructures
    E.g. Herringbone mixers
  LNP structure dependent on lipid and nucleic acid concentration at input
    E.g. lipid-in-ethanol solution
  Some literature regarding scalable and precise pharma manufacturing
    Challenging!
Notes: LNP
https://www.liposomes.ca/publications/2010s/Leung%20et%20al%202015%20-%20Microfluidic%20Mixing%20-%20A%20General%20Method%20for%20Encapsulating%20Macromolecules%20in%20Lipid%20Nanoparticle%20Systems.pdf
https://www.nature.com/articles/s41598-022-13112-5
https://journals.plos.org/plosone/article?id=10.1371/journal.pone.0187962
https://academic.oup.com/jpp/article/76/12/1572/7816331

## Slide 10: LNP Mixing
Notes: https://www.nature.com/articles/s41598-022-13112-5
https://journals.plos.org/plosone/article?id=10.1371/journal.pone.0187962
https://academic.oup.com/jpp/article/76/12/1572/7816331

## Slide 11
Notes: https://europepmc.org/article/pmc/pmc8752123

## Slide 12: Materials
What factors go into material selection
  Chemical and biologic, and thermal compatibility with intendent reagents, products, and reaction type.
  Optical properties
  Mechanical properties
  Cost, Scalability, and ease of fabrication

## Slide 13: Materials
Materials | Major Properties | Applications / Notes | Basic Fabrication Method | Potential Pain Points
Polydimethylsiloxane (PDMS) | Flexible, transparent, gas permeable, biocompatible | Rapid prototyping, biological assays | Soft/Photo lithography | Delicate and chemically permeable
Glass | Chemically resistant, optically transparent | High-pressure and chemical applications | Soft lithography, CNC/micromachining | Microfabrication difficulties, cost, labour intensive
Silicone | Mechanically robust, precise fabrication | Early microfluidic & high performance devices | 3D printing, soft lithography | Opacity renders it useless for optical applications
Thermoplastics (PMMA, COC, COP, PC) | Inexpensive, mass manufacturing, variable transparency | Disposable devices, commercial chips | Embossing, moulding, 3D printing, laser CNC/micromachining | Some manufacturing methods aren’t great for mass production
Paper | Wicking via capillary action, cheap, disposable | Point-of-care, resource-limited diagnostics | Other | Limited use case
Metals | High durability, temperature and pressure tolerance | Specialised applications & High throughput | CNC/micromachining | Opaque, microfabrication difficulties and cost

## Slide 14: Design & Structure
 | POC | IVD | LNP
Channel Dimensions | 20-500 µm width; 10-300 µm height | 50-300 µm width; 20-200 µm height | 100-700 µm width; 50-200 µm height
Device size | Approx. 26x59mm, 2.2mm thickness | Approx. 70x15 mm | Approx. 62x41mm
Critical feature tolerances | +/- 10-20 µm | +/- 15 µm | +/- 5-10 µm
Flow Rates | n/a | n/a | 10-500 mL/min
Common materials | PDMS (R&D/Research), thermoplastics (commercial) | Thermoplastics (chemical resistance) | Thermoplastics
Common manufacturing | n/a | High-volume via injection moulding | n/a
Other details | … | … | Low Reynolds number

## Slide 15: Design & Structure
Conclusion (1): Channel Dimensions & Device Size dependent on microfluidic application
Conclusion (2): Manufacturing tolerance should be 5-20 µm
Conclusion (3): LNP surface finish should ensure low Reynolds number for non-time-dependent and chaotic mixing

## Slide 16: Manufacturing
Main purpose:
  One off/Accuracy
  Scalability/batch … INC: “Industrialisation and automation”
    Paralysation
Manufacturing Methods
  Most Common 
  Direct Fabrication
  Replication-Based
Method | Materials | Feature Resolution (µm)
Soft Lithography | PDMS, glass, thermoplastics | 50-100 nm
Photo Lithography | PDMS, glass, thermoplastics | 50-100 nm
3D Printing/Additive Manufacturing | Plastics, silicone | 30-1000
Notes: 3D printing - https://www.sciencedirect.com/science/article/pii/S02641275230
Soft Litography - Soft Lithography for Microfluidics: a Review
Diagram - https://www.researchgate.net/publication/370866004_Microfluidic_Device_for_the_Manipulation_of_Microparticles_for_Flow_Cytometry_Application

## Slide 18: Manufacturing – Replication Based Methods
Method | Scale | Materials | Feature Resolution (µm) | Main Use
Casting | Low/Medium | PDMS, elastomers | Dependent | Rapid prototyping
Lamination | Medium | Polymers, adhesives | 50-500 | Multilayer chips
Hot embossing | High | Thermoplastics | 1-100 | Mass production
Injection Moulding | High | Thermoplastics | 1-100 | Commercial batches
Film or Roll-to-Roll | High | Polymers, films | 1-100 | Disposable batches
Notes: https://pmc.ncbi.nlm.nih.gov/articles/PMC8002879/
https://www.mdpi.com/2411-5134/3/3/60
https://pmc.ncbi.nlm.nih.gov/articles/PMC7921936/
https://pmc.ncbi.nlm.nih.gov/articles/PMC9147778/
https://www.nature.com/articles/s41598-024-80332-2

## Slide 19
Casting
Lamination
Hot embossing
Injection Moulding
Film or Roll-to-Roll

## Slide 20: Manufacturing – Direct Fabrication Method
Method | Scale | Materials | Feature Resolution (µm) | Main Use
Micro-Cutting (Mechanical Micromachining) | Low/Medium | Polymer, metal | 50-100 | Prototyping, masters
Ultrasonic Machining | Low | Ceramics, glass | 10-100 | Special substrates
Electro discharge Machining (EDM) | Low/Medium | Metals, conductive | 10-50 | Metal masters
Electrochemical Machining (ECM) | Low/Medium | Metals, conductive | 10-100 | Intricate features
Laser Ablation | Low | Polymers, glass | 1-50 | Fast Prototyping
Electron Beam (E-Beam) Machining | Low | Resists, metals | 10 | Master structures
Focused Ion Beam | Low | Polymers, metals | <1 | Nano masters
Notes: https://pmc.ncbi.nlm.nih.gov/articles/PMC8002879/
https://www.mdpi.com/2411-5134/3/3/60
https://pmc.ncbi.nlm.nih.gov/articles/PMC7921936/
https://pmc.ncbi.nlm.nih.gov/articles/PMC9147778/
https://www.nature.com/articles/s41598-024-80332-2

## Slide 21
Micro-Cutting (Mechanical Micromachining)
Ultrasonic Machining
Electro discharge Machining (EDM)
Electrochemical Machining (ECM)
Laser Ablation
Electron Beam (E-Beam) Machining
Focused Ion Beam

## Slide 22: Roll-to-Roll vs. Injection Moulding

## Slide 23: Roll-to-Roll
R-t-R Process
  Liquid UV curable resin is applied to a polymer substrate (PET, Kema foil HSPL80)
  A cylindrical imprint stamp continuously presses into the resin layer
  UV exposure through PET initiates curing, replicating the stamp pattern
  Produces microfluidic channels continuously
Notes: Kilas reference: R2R: MDPI and ACS Style
Smolka, M.; Ruttloff, S.; Nees, D.; Prietl, C.; Satzinger, V.; Lamprecht, B.; Hütter, P.; Hesse, J.; Kokkinis, G.; Kriechhammer, G.; et al. High Throughput Roll-to-Roll Production of Microfluidic Chips. Proceedings 2018, 2, 1054. https://doi.org/10.3390/proceedings2131054

## Slide 24: Injection Moulding
Injection Moulding
  Micro-injection moulding creates high-quality, precise parts
  Hot melt polymer is injected into a mould, cooled and demoulded
  Offers replicability, cost-effectiveness and automation
  Ideal for mass production of disposable products
  Supports a wide range of thermoplastics

## Slide 25: Injection Moulding & R-t-R Comparison
Factor | Injection Moulding | Roll-to-Roll | Lithography
Accuracy | 0.5-5 µm | 10-50 µm | 1-2 µm (best is 100µm)
Throughput | 1000s-100000s | 100,000-continous | <1000 (mainly used for R&D/prototyping)
Material | Most Thermoplastics | Flexible polymers, some thermoplastics | Photoresist, PDMS, glass
Setup Cost | High | Very high, but reduces with per unit cost | Low-medium
Complexity | Good for complex geometries | Challenging for complex 3D geometries | Best for fine details
Notes: https://pmc.ncbi.nlm.nih.gov/articles/PMC8002879/
https://www.mdpi.com/2411-5134/3/3/60
https://www.frontiersin.org/journals/bioengineering-and-biotechnology/articles/10.3389/fbioe.2024.1421831/full
https://www.nature.com/articles/nprot.2009.234
https://www.sciencedirect.com/science/article/pii/S0168365922000918
https://link.springer.com/article/10.1007/s44371-025-00133-y

## Slide 26: Feasibility and Best Use Cases
Injection Moulding
  Highly feasible for commercial, high-volume production
  Best accuracy and broad material options
  Requires significant upfront investment
Roll-to-Roll
  Excellent for low-cost, mass production (e.g. disposable)
  Moderate accuracy, limited to certain designs and flexible materials
  Ideal for paper or film based diagnostic strips (and sometimes electronics!)
Lithography (Soft/Photo)
  Gold standard for research, prototyping and designs requiring sub-micron precision
  Limited scalability for routine high-volume manufacturing

## Slide 27: Recommendations
LNP pharmaceutical chips
  Use injection moulding for commercial scale & lithography for R&D
(Disposable) Diagnostics
  Roll-to-Roll for mass-market, cost-sensitive use
Conclusion
  Manufacturing method is dependent on microfluidic use case (POC/IVD/LNP), design and scaling method
  Further research needs to be done into method of scaling up
    E.g. using chip, channel paralysation or 3rd party flow rate machine

## Slide 28: Appendix